# Instruction manual antenna and cable test set

## Contents

Instruction manual antenna and cable test set	1

Contents	2

General Description	3

Introduction	3

Technical data	3

General	4

Operating Instructions	5

Operating Instructions	5

Initial Instructions	5

VSWR	5

Correcting VSWR for Line Loss	6

Antenna Impedance	7

Precautions	7

Calibration	7

VSWR	7

Some useful formulae relating to VSWR	8

Antenna Voltage Standing Wave Ratio (VSWR)	8

VSWR influence on potential range	8

Cable Loss influence on potential range	9

Total System performance	9

Cable loss influence on VSWR	10

Measuring prototype antennas	11

Measuring “Rubber Duck” antennas	12

## General Description

### Introduction

This test set is designed to measure accurately the VSWR of antennas or other RF loads referred to 50 Ω. It will also determine feeder loss by measuring the effective mismatch of a cable when the far end is in a short circuit.

At frequencies below 150MHz, the Antenna Impedance Z can be determined. The test set is used in combination with an ICOM IC-Q7 VHF/UHF handheld radio, or any other handheld radio capable of delivering 200 to 500mW of RF power over the desired range. The frequency range for the IC-Q7 is 136-174MHz (VHF) and 400-470MHz (UHF).

There is one analogue readout, displaying the measuring bridge unbalance voltage.

There is one control, to set the meter to "Full scale" at the measurement frequency.

The meter requires no electrical power.

### Technical data

Frequency coverage

Bridge: 30MHz to 500MHz

IC-Q7e: 136-174MHz and 400-470MHz

Other radios may be used with the bridge to do measurements outside the IC-Q7's frequency range

Impedance

50 Ω

RF level at measurement connector

6dB below RF input level

17dBm with an input of 23dBm from IC-Q7e

Measurement connector

SMA type socket

### General

Figure "RF Bridge Circuit Diagram"

Uref = (RFin/2)

Uswr = (RFin/2) * (Zx-50)/(Zx+50)

Uz = RFin * (Zx/(Zx+50))

Where RFin is the peak value of the RF input signal.

Figure "Meter Circuit Diagram"

The test set utilises an RF bridge circuit to compare the device under test with a precision 50Ω resistor as shown in Figure "RF Bridge Circuit Diagram".

The DC Output voltages of this bridge are connected to an analogue instrument through a variable voltage divider.

Note that when Zx = R = 50Ω, there is no Uswr output from the bridge and as Zx starts to differ from R, the output (DC) increases and finally when Zx is open or short circuit, reaches a maximum of RFin / 2. This output voltage is proportional to the reflection coefficient of the device under test and after suitable processing may be displayed on as SWR on a meter.

To realise the maximum accuracy of the bridge, great care is required in its assembly.

## Operating Instructions

### Initial Instructions

Connect the Radio (IC-Q7 or other) to the SMA connector RF in on the RF bridge.

### VSWR

- Set up the radio to transmit at the desired operating frequency.

- Attach the device under test to the connector using, if necessary, a short piece of suitable coaxial cable and connectors.

- Place the Set/Measure switch in the Set position.

- Place the SWR/Impedance switch in the SWR position.

- Apply RF power to the bridge and set the meter to read full scale (S) using the Sensitivity control.

- Place the Set/Measure switch in the Measure position.

- Read the VSWR from the meter. The table below converts the VSWR to Return Loss in dB for your convenience.

| VSWR | Return Loss (dB) |
| --- | --- |
| 1 |  |
| 1,1 | -26,4 |
| 1,2 | -20,8 |
| 1,3 | -17,7 |
| 1,4 | -15,6 |
| 1,5 | -14,0 |
| 1,6 | -12,7 |
| 1,7 | -11,7 |
| 1,8 | -10,9 |
| 1,9 | -10,2 |
| 2 | -9,5 |
| 2,5 | -7,4 |
| 3 | -6,0 |
| 4 | -4,4 |
| 5 | -3,5 |
| 7 | -2,5 |

### Line Loss

- Set up the radio to transmit at the desired operating frequency.

- Attach the co-axial line to the connector using, if necessary, a suitable adapter.

- Place the Set/Measure switch in the Set position.

- Place the SWR/Impedance switch in the SWR position.

- Apply RF power to the bridge and set the meter to read full scale (S) using the Sensitivity control.

- Place the Set/Measure switch in the Measure position.

- Read the VSWR from the meter.

- Determine the line loss from the meter reading using the table:

| VSWR | Loss |
| --- | --- |
| 1 |  |
| 1,1 | -13,2 |
| 1,2 | -10,4 |
| 1,3 | -8,8 |
| 1,4 | -7,8 |
| 1,5 | -7,0 |
| 1,6 | -6,4 |
| 1,7 | -5,9 |
| 1,8 | -5,4 |
| 1,9 | -5,1 |
| 2 | -4,8 |
| 2,5 | -3,7 |
| 3 | -3,0 |
| 4 | -2,2 |
| 5 | -1,8 |
| 7 | -1,2 |

### Correcting VSWR for Line Loss

In the Figure "Nomograph for correction of VSWR for cable attenuation", plot the Cable attenuation and the measured VSWR. Draw a line through both points and read the Actual VSWR from the third line.

Figure "Nomograph for correction of VSWR for cable attenuation".

### Antenna Impedance

- Set up the radio to transmit at the desired operating frequency.

- Attach the Device Under Test to the connector using, if necessary, a suitable adapter.

- Place the Set/Measure switch in the Set position.

- Place the SWR/Impedance switch in the Impedance position.

- Apply RF power to the bridge and set the meter to read 50Ω using the Sensitivity control.

- Place the Set/Measure switch in the Measure position.

- Read the Impedance from the meter.

Antenna Impedance reading are usable to 150MHz. Above this frequency, no reliable measurement results can be obtained.

Connect the Device Under Test directly to the measurement connector, or through a (multiple of) half wavelength long transmission line (L = n * ½ λ). Use of other lengths of transmission line will lead to incorrect measurement results, due to impedance transformation by the transmission line.

Use the antenna impedance to find the resonance frequency of an antenna (minimum Z).

### Precautions

Certain precautions should be observed to obtain accurate results on antenna systems:

- It is essential when measuring an antenna, as distinct from a system, to attach the supplied cable direct to the base connector of the antenna. This ensures the highest accuracy by eliminating the effect of long lossy cables between the antenna and test set. The attenuation for this cable is shown in Figure "Attenuation of Test Cable".

- When needing to know the loss of a co-axial cable accurately, it is necessary to perform two or three measurements not only at the frequency of interest but either side of this frequency to average out any resonant effects on the cable.

- When making any measurements, especially at UHF, it is necessary to minimise use of co-axial adapters, as these items can introduce a large degree of uncertainty in any VSWR measurement

- It is inadvisable to operate the test set in the vicinity of high powered transmitters as enough power may enter the test set, via the antenna under test, to damage the diode inside the RF bridge.

## Calibration

### VSWR

The accuracy of the VSWR meter scale should be checked every 12 months by attaching a known calibrated mismatch to the test set. If the reading is not within specification, the instrument must be checked and re-calibrated.

### Some useful formulae relating to VSWR

Antenna Theory

### Antenna Voltage Standing Wave Ratio (VSWR)

Of all the antenna parameters, the VSWR is probably the one most readily measured. Indicating the ratio of antenna impedance to the characteristic impedance of the co-axial feeder, it is possible to use the VSWR figure to calculate the percentage transfer of power to the antenna. If it can be assumed that the antenna meets the necessary mechanical and environmental requirements, the polarisation is correct and the design such that the required coverage can be achieved, then a knowledge of the antenna VSWR across the frequency band of interest will enable performance calculations to be made.

### VSWR influence on potential range

Table 1 shows the power transferred to the antenna for various values of VSWR, the power loss in dBs and the percentage loss in potential system operating range. From the table it can be seen that for a VSWR of 2:1, the percentage of power transferred to the antenna is 89% and the reduction in potential range against a perfect match of 1:1 is approximately 6%. If the VSWR increases to 1:5, then the reductions in range potential will rise to approximately 25%. These figures emphasise the importance of maintaining the antenna VSWR at the initial system design value.

Table 1

| VSWR | % Power to antenna | Power loss in dBs | % loss in potential range |
| --- | --- | --- | --- |
| 1 | 100 | 0 | 0 |
| 1.5 | 96 | 0.18 | 2.02 |
| 2 | 89 | 0.51 | 5.7 |
| 2.5 | 82 | 0.88 | 9.65 |
| 3 | 75 | 1.25 | 13.4 |
| 3.5 | 69 | 1.6 | 16.85 |
| 4 | 64 | 1.94 | 20.00 |
| 4.5 | 60 | 2.25 | 22.86 |
| 5 | 56 | 2.55 | 25.46 |
| 6 | 49 | 3.1 | 30.00 |
| 7 | 44 | 3.59 | 33.86 |
| 8 | 40 | 4.03 | 37.15 |
| 9 | 36 | 4.44 | 40.00 |
| 10 | 33 | 4.81 | 42.5 |

### Cable Loss influence on potential range

By performing a cable attenuation measurement, or from the manufacturers specification, it is possible to obtain a nominal value of attenuation for any cable length and then obtain a figure for the reduction in range potential from Figure "% Reduction in range" below.

Figure "% Reduction in range".

### Total System performance

The effect of the combination of VSWR and cable attenuation on a system performance can be determined by combining the effect of VSWR from table 1 and cable attenuation.

For example, a VSWR of 1: 2.5 gives a power loss of 0.88dB. Combined with a cable loss of 1.5 dB, the reduction in potential range will be 25% according to the table.

### Cable loss influence on VSWR

Apart from its effect on range potential, the attenuation of the co-axial feeder can loss mask the true VSWR of the antenna to which it is connected. An extreme example of this effect is shown in Figure "VSWR masked by Cable attenuation", which gives the apparent VSWR of a short circuited cable against increasing cable attenuation. This is an extreme case, but with some Rx only antennas having a VSWR of 5:1, it can be seen from this figure, that 10 meters of RG-58C/U would give sufficient attenuation at 145 MHz to make a short circuited antenna appear to be acceptable.

Figure "VSWR masked by Cable attenuation".

### Measuring prototype antennas

When building and measuring a prototype antenna, you may encounter problems originating from radiation coming from measurement equipment and cables.

The problems caused by resonating equipment and/ or cables may include:

Change in input impedance at resonance

Change in frequency of resonance.

New resonances appearing.

Change in efficiency.

Change in radiation pattern.

Equipment breaking from the power put into it, especially other electronics

Obscure things like excess distortion.

The problem of changing the pattern is by far the hardest to crack.

The problem is different in prototyping to in final, installed use. By installation I mean installation in its most general term, not just hams installing low-frequency antennas close to the ground, also installation of high-frequency antennas on equipment.  There are two approaches to fixing this problem in installed antennas.

Install the antenna, taking into account the radiation due to cables, masts, things near the installation etc.  Leave the antenna unbalanced and use it.

Balance the antenna well, so the environment can be ignored to a certain extent.

In prototyping there are two approaches also.

Prototype in the place where the installed antenna will go.  Build what works best there.

Balance the antenna in prototyping, then later balance it in the installed antenna.

Other approaches may work -such as constructing an environment similar to the environment that the antenna will be installed in, but can be less reliable.

Designers choose both options. In general antenna designers need to know if their antennas are radiating or not radiating because of the items they are connected to. One way is to balance the antenna with a BALUN and observe. This is a good approach. On the internet and elsewhere, articles can be found on BALUN design. But, without building a BALUN you can make some estimates to see if there is a problem.

Let's say you think you have a well balanced antenna. Here are some methods that can tell you if you've got a problem, they don't though tell you if you for certain that you *haven't* got a problem. Apply the methods in the order they are listed:

Look for a series of narrow resonances spaced in regular periods of frequency

If you have a series of resonances like this it is quite likely that you have a lossy transmission line causing them, not the antenna.

Short wire

Suspecting that the antenna in question is balanced, attach a λ/4 wire to the feed-point. If you have incorporated a BALUN, place it between the BALUN and the generator. Trail this wire in a different axis to that of the antenna. e.g. if the antenna is in XY plane, align it to YZ or XZ plane.

If the resonance moves in frequency you may have a problem. After trying λ /4 try smaller lengths, if these still have an effect then the cable current is large.

Move the items behind the feed through 90 deg

Move the items attached behind the feed or the BALUN through 90 deg into another axis. If the resonances change a lot then you have a problem. If the pattern changes then you can't be sure that it's not just because the items you've moved is not just re-radiating differently in it's new position. You should only try this if you've predicted that your antenna should stay balanced if the feed is moved, many antennas are considered to be balanced (or close to balanced) only if they are fed from a certain direction.

Check the radiation is in the polarisation you expect

This is fairly obvious, if the opposite polarisation is radiating more than the one you think should be radiating then there is probably a problem.

Add a loss mechanism

If you happen to have some Radio Absorbing Material (RAM) knocking around you can do this. If you don't then it's not so easy. Use broadband RAM to do this. Attach the RAM to the feed-point as with the short wire above, around the apparatus attached to the feed. If the bandwidth increases then there is cable radiation, the extra bandwidth is caused by the RAM absorbing the energy. Look for large changes in the bandwidth (like 3dB – 2x). Using ferrites ferrules is an extension of this method.

Simulate it

Simulate the antenna installation including the feed, look at the current on the feed.

Also, try simulating it with the feed removed.

### Measuring “Rubber Duck” antennas

One of the antennas most likely to be measured is the handheld radio’s stock antenna, usually a rubber duck style of antenna.

Handheld radio antennas are usually ¼ wave resonant antennas. This means, that the antenna is only half an antenna. The other half of the antenna is the hand and arm holding the radio. Thus, measuring the stock antenna can result in some erratic measurements. In order to accurately measure a handheld radio antenna, this antenna must be measured against a supplementing set of radials.